## Tabelle1
Lot | Kit code | Quantity | Unit | Descriptor (English) | Specifications | Packing Remarks | item offered by supplier (manufacturer + product name) | unit price | total price | remarks
Lot 1 | PPE-03-01 | 500 | pcs | Face shield, reusable | Lightweight, reusable face shield providing protection against liquid splash and particles and conforming to EC89/686/ECC Category III: EN 166 or equivalent: - Full face-length from headband down by approx. 30 cm - Clear polycarbonate shield - Screen thickness min. 1 mm - Adjustable headband, designed for comfort in tropic climate - Anti-fog coating - Fits easily over prescription glasses or protective eyewear - Resistant to chemicals/highly concentrated chlorine solutions | 50 pcs per box
Lot 2 | PPE-03-02 | 1500 | pcs | Face shield disposable | Lightweight, single use face shield providing protection against liquid splash and particles and conforming to EC89/686/ECC Category III: EN 166 or equivalent: - Full face-length from headband down by approx. 30 cm - Clear screen - Screen thickness approximately 1 mm - Adjustable/elastic headband, designed for comfort in tropic climate - Anti-fog coating - Latex-free - Fits easily over prescription glasses or protective eyewear - SHELF LIFE: 10 YEARS | 100 pcs per box
Lot 3 | PPE-04-01 | 400 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size S | Personal protective equipment that fully covers the wearer's body from head to ankles. Intended to be worn over a surgical tunic and trousers to protect medical and non-medical staff from exposure to inorganic chemicals and infective biological agents (e.g. Ebola). Specifications For reference only: DuPont Tychem C, CHA5®, or alternative/similar Norms Disposable coverall against biological hazards (EN14126) - according to EN13982-1 type 5, EN13034 type 6, EN14605 type 4, personal protection equipment category III according to European legislation or equivalent US legislation Protective clothing (PPE) Category III complex design: Chemical protective clothing Types 3, 4, 5, 6 Protective clothing against infective agents Type 3B (liquid tight) and/or 4B (spray tight) Meets the following European standards requirements: EN 340: 2003 Protective clothing, general requirements Norms applicable to the fabric : EN 368/EN ISO 6530: Resistance of materials to penetration by chemicals liquid EN 14126: 2005. Protection against infective agents from risk groups 1,2,3,4 ISO 16603 : Determination of the resistance of protective clothing materials to penetration by blood and body fluids, test method using synthetic, under hydrostatic = 20 kPa - Class 6/6 ISO 16604 : Determination of resistance of protective clothing materials to penetration by blood-borne pathogens, test method using Phi-X 174 bacteriophage, under hydrostatic = 20 kPa - Class 6/6 ISO 22612 : Test method for resistance to dry microbial (bacteria) penetration = log cfu ≤ 1 - Classe 3 ISO 22611 : Test method for resistance to penetration by biologically contaminated aerosols, using Staphylococcus aureus = log ratio >5 - Class 3/3 ISO 22610 : Test method to determine the resistance to wet bacterial penetration, when subjected to mechanical rubbing = > 75 min - Classe 6/6 EN 14325 : 2004 : Test methods and performance classification of chemical protective clothing materials, seams, joins and assemblages Abrasion resistance: >2000 cycles/Class 6 of 6 Flex cracking resistance: >100 000 cycles/Class 6 of 6 Trapezoidal tear resistance: at least Class 3 of 6 Tensile strength (max. tear): Class 2 of 6 Puncture resistance: Class 2 of 6 Resistance to ignition: at least Class 1 of 3 Norms applicable to the whole coverall EN 1073-2 : 2002 : Requirements and test methods for non-ventilated protective clothing against particulate radioactive contamination = Class 1/3 EN 1149-5 & EN 1149-1 : Electrostatic properties = Class 2 EN ISO 13935-2 : 2004 : Determination of maximum force to seam rupture using the grab method = >125N, level 4/6 EN 14605 : 2005 : Protective clothing Types 3-4: Protection against liquid chemicals (spray-tight) ISO 17491-3 : 2008 : Test methods for clothing providing protection against chemicals - Part 3: Determination of resistance to penetration by a jet of liquid (jet test) = pass ISO 17491-4 : 2008 : Test methods for clothing providing protection against chemicals - Part 4: Determination of resistance to penetration by a spray of liquid (spray test) = pass EN ISO 13982-1 : 2004 : Performance requirements for chemical protective clothing providing protection to the full body against airborne solid particulates (type 5 clothing) = pass EN 13034 : 2005 : Performance requirements for chemical protective clothing offering limited protective performance against liquid chemicals (type 6 clothing) = pass Technical specifications Substrate of Tyvek® fabric (spunbonded olefin sheet made of high density polyethylene fibres) with additional polymeric coating Particle tight and waterproof fabric Stitched and taped seams for better seam strength and to makes it completely leak-proof and waterproof Elastic thumb-loops latex-free attached to the end of the sleeves Zipper: nylon on polyester braid with self-adhesive protective flap Lint-free, latex-free, halogen compounds-free Elastic integrated hood with self-adhesive chin flap Elastic cuffs and ankles Elastic waist With loops for thumb Thickness: 180 μm Colour: yellow Unit presentation, non sterile, for single use | 25pcs per box
| PPE-04-02 | 800 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size M |  | 25pcs per box
| PPE-04-03 | 800 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size L |  | 25pcs per box
| PPE-04-04 | 500 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size XL |  | 25pcs per box
| PPE-04-05 | 200 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size 2XL |  | 25pcs per box
| PPE-04-06 | 200 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size 3XL |  | 25pcs per box
| PPE-04-07 | 100 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size 4XL |  | 20pcs per box
| PPE-04-08 | 100 | pcs | COVERALL, PROTECTIVE, Tychem C or alternative, hooded, yellow, size 5XL |  | 20pcs per box
Lot 4 | PPE-05-01 | 400 | pcs | Scrubs Unisex, set, TROUSERS and TUNIC, Disposable, non-sterile, woven, blue, size S | Disposable scrub pants are available with a drawstring waist or elastic waist. Disposable scrub Shirts with short sleeve, round neck and front pocket Made from breathable, fluid resistant multi-layer lightweight material. Pants have one rear pocket. Fabric Type Breathable Mutli-Layer Fluid R HPIS CLASSIFICATION 760_20_110_10 Latex Free Yes Length Regular MATERIAL Medium Weight Multi-Ply | 200 pcs per box
| PPE-05-02 | 1000 | pcs | Scrubs Unisex, set, TROUSERS and TUNIC, Disposable, non-sterile, woven, blue, size M |  | 200 pcs per box
| PPE-05-02 | 1000 | pcs | Scrubs Unisex, set, TROUSERS and TUNIC, Disposable, non-sterile, woven, blue, size L |  | 200 pcs per box
| PPE-05-03 | 1000 | pcs | Scrubs Unisex, set, TROUSERS and TUNIC, Disposable, non-sterile, woven, blue, size XL |  | 200 pcs per box
| PPE-05-03 | 600 | pcs | Scrubs Unisex, set, TROUSERS and TUNIC, Disposable, non-sterile, woven, blue, size XXL |  | 200 pcs per box
| PPE-05-04 | 400 | pcs | Scrubs Unisex, set, TROUSERS and TUNIC, Disposable, non-sterile, woven, blue, size 4XL |  | 200 pcs per box
Lot 5 | PPE-06-01 | 2000 | pcs | GOWN, isolation, non-sterile, nonwoven, yellow, disposable size M | Protective clothing worn over a scrub shirt/trouser in order to have protection for moderate contact with fluids Specifications/norms according to DIN EN 14126, ISO22610 and ISO22612 Technical specifications: Nonwoven (SMS or SMMS [Spunbond-Meltblown-Spunbond]) Impervious supplementary protection layer (polyethylene film, nonwoven, etc.) in the critical zones (chest area and forearms) Wraparound (not over the head) Long sleeves Cuffs made of elastic jersey Latex-free Dimensions: Comes in 3 sizes: M/Universal, Large and XLarge Length (shoulder seam to mid calf): size M: ± 110 cm size L: ± 136 cm size XLarge: ± 145 cm | 100 pcs per box
| PPE-06-01 | 1500 | pcs | GOWN, isolation, non-sterile, nonwoven, yellow, disposable size L |  | 100 pcs per box
| PPE-06-01 | 500 | pcs | GOWN, isolation, non-sterile, nonwoven, yellow, disposable size XL |  | 100 pcs per box
| PPE-06-02 | 1000 | pcs | GOWN, isolation, non-sterile, nonwoven, blue, disposable size M | Protective clothing worn over a scrub shirt/trouser in order to have protection for moderate contact with fluids Specifications/norms according to DIN EN 14126, ISO22610 and ISO22612 Technical specifications: Nonwoven (SMS or SMMS [Spunbond-Meltblown-Spunbond]) Impervious supplementary protection layer (polyethylene film, nonwoven, etc.) in the critical zones (chest area and forearms) Wraparound (not over the head) Long sleeves Cuffs made of elastic jersey Latex-free Dimensions: Comes in 3 sizes: M/Universal, Large and XLarge Length (shoulder seam to mid calf): size M: ± 110 cm size L: ± 136 cm size XLarge: ± 145 cm | 100 pcs per box
| PPE-06-02 | 700 | pcs | GOWN, isolation, non-sterile, nonwoven, blue, disposable size L |  | 100 pcs per box
| PPE-06-02 | 300 | pcs | GOWN, isolation, non-sterile, nonwoven,blue, disposable size XL |  | 100 pcs per box
| PPE-06-03 | 500 | pcs | GOWN, isolation, non-sterile, nonwoven,white, disposable size M | Protective clothing worn over a scrub shirt/trouser in order to have protection for moderate contact with fluids Specifications/norms according to DIN EN 14126, ISO22610 and ISO22612 Technical specifications: Nonwoven (SMS or SMMS [Spunbond-Meltblown-Spunbond]) Impervious supplementary protection layer (polyethylene film, nonwoven, etc.) in the critical zones (chest area and forearms) Wraparound (not over the head) Long sleeves Cuffs made of elastic jersey Latex-free Dimensions: Comes in 3 sizes: M/Universal, Large and XLarge Length (shoulder seam to mid calf): size M: ± 110 cm size L: ± 136 cm size XLarge: ± 145 cm | 100 pcs per box
| PPE-06-03 | 300 | pcs | GOWN, isolation, non-sterile, nonwoven,white, disposable size L |  | 100 pcs per box
| PPE-06-03 | 200 | pcs | GOWN, isolation, non-sterile, nonwoven,white, disposable size XL |  | 100 pcs per box
Lot 6 | PPE-07-01 | 50 | pcs | BOOTS, rubber, white, pair, size 38 | WORK-boots, white/grey, according to EN347 and ISO6112 - White shaft, grey deep-tread sole (super non-slip) - Shaft height according to size, 30-38 cm - Smooth, pore-free surface for perfect hygiene | 25 pcs per Box
| PPE-07-02 | 50 | pcs | BOOTS, rubber, white, pair, size 39 |  | 25 pcs per Box
| PPE-07-03 | 100 | pcs | BOOTS, rubber, white, pair, size 40 |  | 25 pcs per Box
| PPE-07-04 | 100 | pcs | BOOTS, rubber, white, pair, size 42 |  | 25 pcs per Box
| PPE-07-05 | 100 | pcs | BOOTS, rubber, white, pair, size 44 |  | 25 pcs per Box
| PPE-07-06 | 50 | pcs | BOOTS, rubber, white, pair, size 46 |  | 25 pcs per Box
| PPE-07-07 | 50 | pcs | BOOTS, rubber, white, pair, size 48 |  | 25 pcs per Box
| PPE-07-08 | 50 | pcs | BOOTS, rubber, white, pair, size 52 |  | 25 pcs per Box
| PPE-07-09 | 50 | pcs | BOOTS, rubber, white, pair, size 54 |  | 25 pcs per Box
Lot 7 | PPE-08-01 | 20000 | pcs | Surgical Cap, non-woven, standard size, blue, disposable | Type: Bouffant Fasting type: Elastic Colour: Blue Adult model, standard size, rubber banded Material: Non-woven, disposable Item Weight: 15gm Thickness: 0,05mm Packaging: carton Shelf life up to 10 years from the date of manufactur | 2000pcs per box
Lot 8 | PPE-09-01 | 12000 | pcs | GLOVE, EXAMINATION, nitrile, non sterile, extended cuff, S | Non sterile glove with extended cuff used to protect both patient and staff from infections transmitted through contact with the hands. Nitrile, non-sterile, ambidextrous, disposable; semisolid; high qualit; powder-free; cuff edge straight or rolled; acco rding to EN 455-1-2-3-4/2001-2007-2009, and Directives 93/42/EC class I Specifications Norms • EN 455-1: 2001 "Medical gloves for single use - Requirements and testing for freedom from holes." AQL = 1.5 • EN 455-2: 2009 + A2: 2013 "Medical gloves for single use - Requirements and testing for physical properties" Force at break ≥ 6N before and after ageing. • EN 455-3: 2007 "Medical gloves for single use -Requirements and testing for biological evaluation" • EN 455-4: 2009 "Medical gloves for single use - Part 4: Requirements and testing for shelf life determination" • EN 374-2: 2015 "Protective gloves against dangerous chemicals and microorganisms - Part 2: Determination of resistance to penetration" Edition: 3. AQL = 1.5 Technical specifications • 100% nitrile, non powdered • Ambidextrous shape (fits either hand) • Thickness at fingers: max 0.1 mm • Cuff: extended with straight edge • Colour : blue or purple • Non sterile, for single use Dimensions 5 sizes: small (S), medium (M), large (L), x-large (XL) and xx-large (XXL) Size Minimum length Palm width small 290 mm 85 ± 5 mm medium 290 mm 95 ± 5 mm large 290 mm 105 ± 5 mm x-large 290 mm 115 ± 5 mm xx-large 290 mm > 120 mm Packaging Cardboard distribution box of 100 gloves. Shelf life: 10 years | 100 pcs per box
| PPE-09-02 | 20000 | pcs | GLOVE, EXAMINATION, nitrile, non sterile, extended cuff, M |  | 100 pcs per box
| PPE-09-03 | 20000 | pcs | GLOVE, EXAMINATION, nitrile, non sterile, extended cuff,L |  | 100 pcs per box
| PPE-09-04 | 12000 | pcs | GLOVE, EXAMINATION, nitrile, non sterile, extended cuff, XL |  | 100 pcs per box
| PPE-09-05 | 8000 | pcs | GLOVE, EXAMINATION, nitrile, non sterile, extended cuff, XXL |  | 100 pcs per box
Lot 9 | PPE-10-01 | 1000 | pcs | HOOD, Tyvek (similar or alternative), with mask, non-woven | Must cover shoulders entirely; ties at back of the head, neck and underneath arms Fabric • EN14126 with level ≥4/6 for test ISO 16603 and preferably level ≥4/6 for test ISO 16604 o sewed and covered (=sealed) or welded with ultrasounds Form/design : • Must cover shoulders entirely; ties at back of the head, neck and underneath arms o 3 pairs of straps to tighten for fixation behind the head o 1 or 2 pairs of straps for fixation at the height of the torso o With integrated mask • Non sterile, for single use • Unit presentation | 100 pcs per box
Lot 10 | PPE-11-01 | 60 | pcs | FULL FACE RESPIRATOR MASK with bayonet connections, reusable | Face seal Double-layer face seal with triple sealing action provides reliable protection and a secure fit for various face types. Universal size 5-point harness Quick and easy to don Filter connection: Two side-positioned bayonet connections for Bayonet Filter. Both filters are inserted in a top-down direction. Lens, wide field vision, chemical, thermal, and mechanical resistance. Material: Mask body EPDM with high ageing resistance, extremely hypoallergenic Weight approx. 540 g Approvals EN136 Cl.2 with CE mark and NIOSH Maintenance Spare parts available | 10pcs per box
Lot 11 | PPE-11-03 | 120 | pcs | RESPIRATORY FILTERS, pair, for Full Facepiece Reusable Respirator with bayonet connection | Replaceable respirator filter pair, suitable for repeated use with compatible reusable respirator systems. Connection: Bayonet connection or manufacturer-approved connection, compatible with the specified full-facepiece respirator. Protection: against organic and acidic gases, vapors, sulfur dioxide, and health-hazardous particles Filter Class: A2, P3, A1, B1, E1, K1, Hg. Compatibility: Must be fully compatible with the supplied full-facepiece respirator and any approved filter retainers/adapters. Standard: Compliant with EN 143 and applicable EU PPE requirements, with CE marking where required | 20pcs per box
Lot 12 | PPE-12-01 | 3000 | pcs | RESPIRATOR, FFP2 Mask, folding, duck bill, without valve | The FFP2 mask without valve. This is tested and certified in accordance with the EN 149:2001 + A1:2009 standard and offers reliable protection against particles, aerosols and other harmful substances. The FFP2 mask, which is usually used in the general work environment, is intended to provide protection against bodily fluids, among other things. The FFP2 masks offer self-protection and protection for others. This FFP2 mask offers 5-layer protection and has no valve. The FFP2 mask offers a comfortable fit and has two loops for the ears. Product features of the FFP2 masks • FFP2 mask in accordance with PPE Regulation 2016/425 category 3 • Tested and certified CE2195, formerly CE 2163 according to the standard EN 149:2001 + A1:2009 • 5-layer protection against particles, aerosols and other harmful substances • Hygienic individually made Packaged* (better transport and prevents contamination) • Self-protection and protection of others • Ear loops • Dimensions: width 10.5cm / height 15.5 cm • Color: white • Integrated nose clip (optimal fit) • SMSS – Material • Bung-bound and melt-blown (meltbown) • NR = single use | 25pcs per box
Lot 13 | PPE-12-02 | 1500 | pcs | RESPIRATOR, FFP3 Mask, without valve, duck bill | Filtering device covering nose, mouth and chin, used to protect the wearer against airborne or droplets transmitted infectious agents. Adjustable nose clip with the double adjustable headband High fluid resistance, good breathability, internal and external faces should be clearly identified, structured design that does not collapse against the mouth. FFP3 without valve; safe high-risk-mask, duck bill; particle filtration efficiency PPE (0,1 micron)>99%, bacterial filtration efficiency (0,3 micron)>99% FFP3 must be written on the respirator itself Norm EN 149:2001+A1:2009 | 20pcs per box
| PPE-12-03 | 1500 | pcs | RESPIRATOR , FFP3 Mask, with valve, duck bill | Filtering device covering nose, mouth and chin, used to protect the wearer against airborne or droplets transmitted infectious agents. Adjustable nose clip with the double adjustable headband High fluid resistance, good breathability, internal and external faces should be clearly identified, structured design that does not collapse against the mouth. FFP3 with valve; safe high-risk-mask, duck bill; particle filtration efficiency PPE (0,1 micron)>99%, bacterial filtration efficiency (0,3 micron)>99% FFP3 must be written on the respirator itself Norm EN 149:2001+A1:2009 | 20pcs per box
Lot 14 | PPE-13-01 | 200 | pcs | GOGGLES, SAFETY glasses, regular nose, plastic | Goggles, or safety glasses, protective eyewear to prevent particulates, water or chemicals dust and splashing from striking the eyes deep curved PVC frame, close fit and versality to wear over prescription glasses. Clear polycarbone lenses. Protection against impact from flying particles, droplets, splashes. Vented body (indirect ventilation). Chemicals resistant Lens: Material: mechanically strong but less flexible polycarbonate (PC) or a better chemical resistant acetate type (PA) are equivalent options Minimum mechanical resistance (resistance to shocks of low-energy particles): withstands a bead of 6 mm and 0.85 g at 45m/s impact Clear lens for protection against UV light Panoramic view 180° The optical zone for every eye must be at least 32 mm horizontally and 25 mm vertically Replaceable Treated against fogging and against scratching Must be easy to disinfect and resistant to chlorine 0.5% solution Frame: Good seal with the skin of the face Shape of nose bridge: Regular Parts that touch the skin may not be allergenic or irritating Easy to combine with corrective glasses Made in thermoplastic elastomer Flexible frame that easily fits all face contours and accommodates prescription glasses Broad and fully adjustable headband that does not become loose during clinical activity. Must be easy to disinfect and resistant to chlorine 0.5% solution Protection against impact from flying particles, droplets, splashes. Vented body (indirect ventilation). Chemicals resistant | 25pcs per box
| PPE-13-02 | 200 | pcs | GOGGLES, SAFETY glasses, wide nose, plastic |  | 25pcs per box
Lot 15 | PPE-13-01 | 80 | pcs | ANTI-FOG SPRAY, for goggles | Anti-fog spray for goggles Nanotechnology: high-quality sealing/impregnation environmentally friendly: water-based, without toxic/hazardous content | 10pcs per box
Lot 16 | PPE-14-02 | 50 | pcs | GLOVES, PROTECTIVE, nitrile, reusable, pair, 8/ M | Pair of high-comfort and chemical resistant heavy-duty gloves - Cotton flocking, reducing sweat for the wearer - Fully reusable, with an unequaled abrasion resistance - Provides superb comfort for the wearer - Reversed lozenge finish further enhancing levels of grip • High performance nitrile compound, providingan outstanding combination of chemical resistance and strength for optimal results in wet or dry work environments • The glove offers a far superior snag, puncture and abrasion protection compared with rubber or neoprene gloves • High levels of flexibility, comfort and dexterity • Will not swell, weaken or degrade, and does not promote contact dermatitis • Suitable for heavy-duty cleaning applications Standards/Certifications: - EN 420:2003 + A1:2009; 0493 - conforms to EC 89/686/EEC, Category III: EN 374, -2, 3, 4, EN 420 and EN 388 (High cracking-, puncture-and abrasion resistant, levels: ≥ 4-1-0-1) or equivalent standards Material: - Nitrile - Minimum length: 380 mm - Left/right pair - Knit inner lining to facilitate slide-in and removal - Thickness: ≥ 0,425 mm - Cleanable with water and disinfectant (resisting both ethanol solutions 70% and chlorine solutions >0.5%) | 10pcs per box
| PPE-14-03 | 50 | pcs | GLOVES, PROTECTIVE, nitrile, reusable, pair, 9/L |  | 10pcs per box
| PPE-14-04 | 50 | pcs | GLOVES, PROTECTIVE, nitrile, reusable, pair, 10/XL |  | 10pcs per box
| PPE-14-05 | 40 | pcs | GLOVES, PROTECTIVE, nitrile, reusable, pair, 11/XXL |  | 10pcs per box
Lot 17 | PPE-15-01 | 150 | pcs | APRON, reusable, heavy duty, acid resistant, PVC | Worn over work wear when carrying out certain surgical procedures and cleaning work in the sterile services department Material: PVC or nitric rubber Straight apron with bib, back fastening and adjustable neck strap High tear resistance and tear propagation resistance,very flexible and lightweight Welded-on plastic eyes for neck bands and side bands of polyester. Standard adult size approx 120 x 90 cm Thickness: min. 0.3 mm High resistance to fats, acids and heat Can be washed and disinfected Non-sterile, multiple use Approved as clothing for protection against chemicals based on EN 14605 (type 3 and 4) | 25pcs per box
Lot 18 | PPE-16-01 | 1000 | pcs | SURGICAL GLOVES, latex-free, sterile, pair, size 6,5 | One pair of slightly powdered gloves: 1 right-handed, 1 left-handed Material: latex-free Powder: maize starch Physical property: - Waterproof - Stretch proof Minimum force at break for surgical gloves: - before accelerated ageing 10.5N - after accelerated ageing 7.5N Straight sleeved with reinforced hem (rolled or ending in a reinforced band) Moulded fabrication Suppleness: closely fits the morphology of the hand and minimally impairs the wearer’s sense of touch Shape of the glove faithfully accommodates the anatomy of the hand, the thumb offset from the palm and set forward of the index finger Slightly powdered: the interior surfaces of the gloves must be finely coated with powder for easy wearing Low residual protein level (EN 455-3) Recommended wall thickness: - Finger area 0.38 mm - Centre palm area: 0.24 - 0.30 mm / 25 mm - From cuff end: 0.26 mm Size: minimum dimensions recommended for surgical gloves size 6.5 Length: 280 mm Width: 85 ± 5 mm The surgical glove cuff should have a sufficient length to cover adequately the surgeon/nurse wrist when wearing operating theatre gown Initial sterilisation method: Ethylene oxide gas Packaging: - One (1) pair or surgical gloves in an individual sterilised peel pack - Gloves must be supplied with double packaging: one interior layer plus sterilised peel-pack - Gloves must be packed in the interior layer: right and left hand with the sleeves reversed up to the thumb - Labeled with: size, right hand, left hand and position of the sleeves - 50 units (pairs of surgical gloves in peel pack) in protective carton box Labeling: Peel pack and box should mention: name and/or trademark of the manufacturer, sterile, lot number, expiry date by year and month, "for single use". CE + number, product description Extra information required: number of units per secondary packaging Other requirements: conforms to European Norm EN 455 (Parts 1, 2 and 3) | 50pcs per box
| PPE-16-02 | 2500 | pcs | SURGICAL GLOVES, latex-free, sterile, pair, size 7,5 |  | 50pcs per box
| PPE-16-04 | 1000 | pcs | SURGICAL GLOVES, latex-free, u.u., sterile, pair 8,5 |  | 50pcs per box